## Slide 1: Pacemakers and Implantable Cardiac Defibrillators: Software Radio Attacks and Zero-Power Defenses Daniel Halperin, Thomas S. Heydt-Benjamin, Benjamin Ransford, Shane S. Clark, Benessa Defend, Will Morgan Kevin Fu, Tadayoshi Kohno, William H. Maisel
Sangwon Lee
2019. 12. 05
Notes: 심박동기 삽입형 제세동기

## Slide 2: Contents
Introduction
Vulnerabilities
Intercepting ICD communications
Defenses
Conclusions

## Slide 3: Introduction
Security and privacy vulnerabilities in implantable medical devices (IMD).
What are IMDs?
  Medical devices to replace, support, or enhance an biological structure.
Between 1990-2002: 2.6 million IMDs implanted in US patients
Notes: Deep brain neurostimulators - 일정 주기마다 뇌에 전기 신호 보냄. 파키슨 병의 치료에 사용
cochlear implant - 인공 달팽이관
Gastric Stimulators - helps control the chronic nausea and vomiting
Foot drop - 신경성 족하수(걸음걸이 이상) 치료
Insulin pump

## Slide 4
Notes: Deep brain neurostimulators - 일정 주기마다 뇌에 전기 신호 보냄. 파키슨 병의 치료에 사용
cochlear implant - 인공 달팽이관
Gastric Stimulators - helps control the chronic nausea and vomiting
Foot drop - 신경성 족하수(걸음걸이 이상) 치료
Insulin pump

## Slide 5: Implantable Cardioverter Defibrillators (ICDs)
Implanted inside the patient’s body.
Monitors and responds to heart activity
  Pacing - steady small shocks
  Defibrillation - Single larger shock
Power and connectivity
  No physical external connections
  Non-rechargeable internal batteries
  Designed to last for many years

## Slide 6: Implantable Cardioverter Defibrillators (ICDs)
Reprogrammable by the commercial device programmer
  Perform diagnostics
  Read / write patients’ private data
  Adjust therapy settings.
Wireless
Vulnerabilities exist

## Slide 7: Vulnerabilities

## Slide 8: Vulnerabilities
Commercial ICD programmers can be used without authentication.
Wireless communication between the ICD programmer and the ICD is not encrypted. - can be eavesdropped.
ICDs can be re-programmed by an unauthenticated device.

## Slide 9: Insider Attack - Commercial ICD programmer
Programmers can read all ICD information, change all settings.
No technological controls to ensure authorized use.

## Slide 10: Intercepting ICD communications

## Slide 11: Used Equipment for Reverse Engineering
Oscilloscope
  Measures voltage of signal and displays as a waveform
Universal Software Radio Peripheral
  Device that interacts with open source GNU Radio libraries.
Software: GNU Radio, Perl, Matlab
Antennas
USRP
Antennas
ICD
Oscilloscope

## Slide 12: Reverse Engineering Transmissions
Captured RF transmissions using the oscilloscope and USRP at 175 kHz.
Processed captured signals using the softwares to analyze the ICD’s protocols.
ICD
Programmer
OSC
USRP
PC
RF signal capture
Analyze signals

## Slide 13: Reverse Engineering Transmissions
Transmissions from the Programmer
  Capture raw bits between the programmer and the programming head.
  Compare the raw bits with the encoded and modulated RF signals.
Head
Programmer
0
1
1
1
...
0
Compare

## Slide 14: Reverse Engineering Transmissions
Transmissions from ICD
  Made dummy patient name.
  Analyzed RF signal to determine the modulation and encoding scheme.
ICD
Programmer
Patient’s name
AAA
Analyze repeated patterned signals
Transmit information
containing the name

## Slide 15: Eavesdropping
Using GNU Radio library + USRP
  Builds a GNU Radio program to gather data
  Removed 87 lines of C++ code, added 44 lines of code.
Establishing a transaction timeline

## Slide 16: Eavesdropping
Intercepting Patient Data
  Not cryptographically protected
  Cleartext representations of patient data
  Easily extractable
  Personal data
    Patient’s name, date of birth, medical ID number, and patient history.
    Name and phone number of the treating physician
Intercepting Telemetry
  Activated with a magnet within 5cm of the ICD
  Broadcast patient’s electrocardiogram in cleartext
Notes: electrocardiogram - 심전도

## Slide 17: Active Attacks
All active attacks are replay attacks
  Transmit-only
Replay attacks with GNU radio
  Disclosing patient data
  Disclosing cardiac data
  Changing patient name
  Setting the ICD’s clock
  Changing therapies
  Inducing fibrillation

## Slide 18: Changing therapies
Therapies
  ICD’s responses to cardiac events
Replay attack can turns off therapies
After 24 replay attempts, more than one succeeded at disabling all therapies

## Slide 19: Inducing fibrillation
The ICD can induce ventricular fibrillation
  Sends command shock to the patient’s heart, averaging 137.7 V

## Slide 20: Power denial of service attack
Continuous engagement in wireless communication
  Decreases battery life faster

## Slide 21: Defenses

## Slide 22: Defense Goals
Security events should be detectable by patients
Prevent attacks from insiders and outsiders
Draw no power from primary battery

## Slide 23: Zero Power Defense
WISPer - Wireless Identification and Sensing Platform
  Embedded system : RFID circuit + microcontroller
WISPer harvests RF energy from a RFID reader
Three applications:
  Notification
  Authentication
  Sensible key exchange

## Slide 24: Defense 1 - Notification for Patients
When WISPer is activated, beep via piezoelectric speaker
  After beep, notify ICD it can start using radio
Patient aware when ICD is being programmed
Tested with bacon
  Measured 84 dB of sound at the surface of the tissue
    Normal conversation: about 60 dB

## Slide 25: Defense 2 - Authentication
RC5-32/12/16-based Authentication
Only if authentication is successful will ICD be activated
No power is used until authentication succeeds.

## Slide 26: Defense 3 - Sensible Key Exchange
Use audio as a channel for crypto key exchange.
Modulate sound wave using same scheme as radio.
Audible to patient, hard to hear at a distance.
Also uses no power.

## Slide 27: Conclusion
ICD is susceptible to malicious attacks
  From privacy information to intended malfunctioning
Balance between security, privacy, safety, and effectiveness should be achieved
More security should be considered for next-generation ICDs

## Slide 28: Related works
Halperin, Daniel, et al. @ 2008
Security and privacy for implantable medical devices.
→ Earlier work surveys a wide range of IMD security issues, including the need
to balance IMD security and privacy goals with safety and effectiveness goals

## Slide 29: Related works
A. Perrig et Al. @ 2002
SPINS: Security protocols for sensor networks
C. Karlof et Al. @ 2004 TinySec: A link layer security architecture for wireless sensor networks Studied wireless security in low-power environments, especially in the areas of sensor networks
S. Warren Et Al. @ 2005 Interoperability and security in wireless body area network infrastructures Wireless body area networks
→ Proposed zero-power security approaches to eliminate the stored-energy overhead of cryptography

## Slide 30: Related works
D. Spill et Al. @ 2007 BlueSniff: Eve meets Alice and Bluetooth
T. S. Heydt-Benjamin et Al. @ 2007 Vulnerabilities in first-generation RFID-enabled credit cards
→ Demonstrated the utility of software radios by using them to help reverse-engineer and then participate in unknown radio protocols.

## Slide 31: After the paper release

## Slide 32: After the paper release

## Slide 33: After the paper release

## Slide 34
Thank you